# CLEANING THE INCUBATOR

- Remove shelves

- Remove water residues with pipette controller

- Dry the remaining water using sterile cloths

- Spray Virkon or cleaning liquid, and wipe down the incubator surfaces

- Repeat the process with distilled water (spray and wipe)

- Spray Virkon (or cleaning liquid) and wipe down the shelves

- Spray the interior of the incubator with IMS and wipe dry

- Spray the shelves with IMS and carefully reinsert them into the incubator

- Spray IMS and wipe down both sides of the glass door

- On the screen click CONTRA-CON and follow the instructions (takes approx. 25hrs)

- Add 350mL of Distilled water to the water reservoir

- Shut the door to continue the contra-con process

- Put a sign on the incubator; “DO NOT USE!”

- Once the process is complete, remove the residual water and add enough water to cover the water reservoir sensor.

- Click auto-start button and follow the instructions (this could take 10hrs).

- Again, put a sign on the incubator; “DO NOT USE!”

- Once complete, remove the tape and complete the form with date, comments and name.