CLEANING THE INCUBATOR

1. [bookmark: _GoBack]Remove shelves
2. Remove water residues with pipette controller
3. Dry the remaining water using sterile cloths
4. Spray Virkon or cleaning liquid, and wipe down the incubator surfaces
5. Repeat the process with distilled water (spray and wipe)
6. Spray Virkon (or cleaning liquid) and wipe down the shelves
7. Spray the interior of the incubator with IMS and wipe dry
8. Spray the shelves with IMS and carefully reinsert them into the incubator
9. Spray IMS and wipe down both sides of the glass door
10. On the screen click CONTRA-CON and follow the instructions (takes approx. 25hrs)
11. Add 350mL of Distilled water to the water reservoir
12. Shut the door to continue the contra-con process
13. Put a sign on the incubator; “DO NOT USE!”
14. Once the process is complete, remove the residual water and add enough water to cover the water reservoir sensor.
15. Click auto-start button and follow the instructions (this could take 10hrs). 
16. Again, put a sign on the incubator; “DO NOT USE!”
17. Once complete, remove the tape and complete the form with date, comments and name.





